## Chromium WorkFlow
Source | Description | Part # | Storage Conditions | Step(s) in protocol | Why needed
Thermo Fisher Scientific | Qubit® 3.0 Fluorometer** | Q33216 | RT | DNA Quantification | Required for input sample mass quantification. Older version 2.0 is compatible with workflow.
| Qubit dsDNA HS Assay Kit* | Q32854 | RT/4℃ | DNA Quantification | Failure to load the correct mass will adversely impact system performance. Note: 100 assay kit available (Q32851)
| QubitAssay Tubes* | Q32856 | RT | DNA Quantification | Required for proper use of Qubit Fluorometer.
| DynaBeads® MyOne™ Silane Beads* | 37002D | 4°C | GEM Cleanup and Library Construction | Required for Chromium sample clean-up steps. Critical for GEM cleanup.
| Nuclease-Free Water | AM9937 | RT | Reagent Setup and Chromium Instrument Operation | Required for Master Mix and dilution plates for qPCR
Qiagen | Buffer EB* | 19086 | RT | GEM Cleanup and Library Construction | Required for Chromium sample dilutions and clean-up.
Sigma | Ethanol, Pure (200 Proof, anhydrous) | 459836-500ML | RT | GEM Cleanup and Library Construction | Required for Chromium sample clean-up steps
Beckman Coulter | SPRIselect Reagent Kit * | B23318 | RT | GEM Cleanup and Library Construction | Required for proper size selection and cleanup of Chromium libraries
USA Scientific | TempAssure PCR 8-tube strip** | 14024700 | RT | Reagent Setup and Chromium Instrument Operation | Required. The following alternative sources are also compatible: Applied Biosystems ® MicroAmp® 8-tube strip (N8010580) and 8-cap strips (N8010535). Eppendorf 0.2mL PCR tube strips (951010022)
Eppendorf | Eppendorf twin.tecTM 96-Well PCR Plate (Semi-Skirted) * | 951020362 | RT | Chromium Instrument Operation | Required plastics compatible with Chromium partioning oil. Failure to use this plate type will cause GEM partitions to break during thermal cycling. Other twin.tec plates include: 96-Well PCR Plate Divisible Unskirted 003013374 or 96-Well PCR Plate Unskirted 0030133390.
| Eppendorf twin.tecTM 96-Well PCR Plate (Skirted) * | 951020460 | RT | Reagent Setup and Chromium Instrument Operation
| Thermomixer C1.5*** | 5382000023 | RT | GEM Cleanup and Library Construction | Needed for warming Chromium cleanup buffer. Heat block also ok, must go to 70C.
| DNA LoBind Tube Microcentrifuge Tube, 1.5 mL*** | 022431021 | RT | Various Steps | Only use lowbind tubes
| DNA LoBind Tube Microcentrifuge Tube, 2.0 mL*** | 022431048 | RT | Various Steps | Only use lowbind tubes
Rainin | Tips LTS 20UL Filter RT-L10FLR | 17007957 | RT | Various Steps | Recommended for Chromium workflow. Failure to use low retention tips may cause loss of DNA to plastic.
| Tips LTS 200UL Filter RT-L200FLR | 17007961 | RT | Various Steps | Recommended for Chromium workflow. Failure to use low retention tips may cause loss of DNA to plastic.
| Tips LTS 1ML Filter RT-L1000FLR | 17007954 | RT | Various Steps | Recommended for Chromium workflow. Failure to use low retention tips may cause loss of DNA to plastic.
| Wide-orifice LTS presterilized low retention LTS filter tips, 200 µL* | 17014294 | RT | HMW DNA Handling | Wide orifice tips are required to minimize shearing of gDNA during sample handling
| Pipet-Lite LTS Pipette L-2XLS+ | 17014393 | RT | Various Steps | Recommended for Chromium workflow. Required to pipette < 2 uL volumes.
| Pipet-Lite LTS Pipette L-10XLS+ | 17014388 | RT | Various Steps | Optional for lab preference pipetting volumes.
| Pipet-Lite LTS Pipette L-20XLS+ | 17014392 | RT | Various Steps | Recommended for Chromium workflow.
| Pipet-Lite LTS Pipette L-100XLS+ | 17014384 | RT | Various Steps | Optional for lab preference pipetting volumes.
| Pipet-Lite LTS Pipette L-200XLS+ | 17014391 | RT | Various Steps | Recommended for Chromium workflow.
| Pipet-Lite LTS Pipette L-1000XLS+ | 17014382 | RT | Various Steps | Recommended for Chromium workflow.
| Pipet-Lite XLS LTS+ 8-channel, 0.5 - 10 µL | L8-10XLS+ | RT | Reagent Setup and Chromium Instrument Operation | Recommended for Chromium workflow.Recommeded for Chromium chip loading and cleanup steps, processed in batches of 8. Electronic pipettor may be used for mixing steps, lab preference.
| Pipet-Lite Multi Pipette L8-200XLS+ | 17013805 | RT | Reagent Setup and Chromium Instrument Operation | Recommended for Chromium chip loading and cleanup steps, processed in batches of 8. Electronic pipettor may be used for mixing steps, lab preference.
Bio-Rad | C1000 TouchTM Thermalcycler with 96-Deep Well Reaction Module ** | 1851197 | RT | Reagent Setup and Chromium Instrument Operation | Eppendorf MasterCycler® Pro or Thermo Fisher Veriti© 96-Well Thermal Cycler can also be used. MUST BE COMPATIBLE WITH Eppendorf® twin.tec plates and hold a volume of 125uL
| PX1 PCR Plate Sealer*** | 181-4000 | RT | Reagent Setup and Chromium Instrument Operation | Validated for proper heat sealing. Other plate sealers may work. Require 185° C for 6 seconds.
| 8 cap strips | TCS0803 | RT | Reagent Setup and Chromium Instrument Operation | For capping plate when running multiple Chromium chips.
| Piercable Foil Heat Seal** | 181-4040 | RT | Preparing the Plate for Thermalcycling | Validated for proper heat sealing.
| 10% Tween-20 | 166-2404 | RT | DynaBeads Cleanup | Needed for Chromium Cleanup
VWR | Microcentrifuge with 8-Tube Strip Adapter | many*** | RT | Reagent Setup and Chromium Instrument Operation | Needed for spinning tube strips. Recommend access to two.
| Vortex Mixer*** | 10153-838 | RT | Reagent Setup and Chromium Instrument Operation | Required for Chromium Vortex Adaptor compatibility. Cannot guarantee proper mixing of gel beads otherwise. Fits standard vortex attachment head. Recommend access to two.
| Divided Polystyrene Reservoirs*** | 41428-958 | RT | Reagent Setup and Chromium Instrument Operation | Recommended for Chromium workflow post GEM creation. Integra 10 ml Low Residual Volume Disposable Multichannel Pipette Reservoirs are another example.
Agilent | Agilent 2100 Bioanalyzer** |  | RT | Materials – GEM Cleanup and Library Construction | Required for QC after GEM creation and library construction.
| Agilent DNA 1000 Kit** | 5067-1504 | 4°C | Materials – GEM Cleanup and Library Construction | Required for QC after GEM creation and library construction.
| Agilent High Sensitivity DNA Kit** | 5067-4626 | 4°C | Materials – GEM Cleanup and Library Construction | Required for QC after GEM creation and library construction.
General | Tube Racks | -- | -- | Reagent Setup and Chromium Instrument Operation | Various tube volumes
| Gloves | -- | -- | Reagent Setup and Chromium Instrument Operation | For personal protection
| Ice Bucket and Ice | -- | -- | Reagent Setup and Chromium Instrument Operation | For maintaining reagents on ice
| Ricca Chemical Company | 3290-32 | -- | Chip loading | Glycerin (glycerol), 50% (v/v) Aqueous Solution (additional Surrogate Fluid, if needed) Can make your own. See User Guide
| Metal 96 well block | -- | -- | Reagent Setup and Chromium Instrument Operation | For preparing sample Master Mix. Example: CoolRack® XT PCR96
^ Chromium System Training starts with extracted DNA and ends prior to shearing.
^ Note that other items are required for Extraction, and Whole Genome or Whole Exome Sequencing after training. Refer to the Chromium User Guide for complete workflow requirements.
* Items in bold red have been validated by 10X Genomics and are required for Chromium workflow training and system operation. No substitutions are allowed.
** Listed materials are compatible with workflow. Substituting materials may adversely affect system performance and are not supported.
*** Example materials listed. Substitutions may be used.

## DNA Extraction
Source | Description | Part # | Storage Conditions | Step(s) in protocol | Why needed
Qiagen | Qiagen MagAttract HMW DNA Kit** | 67563 | RT/-20°C | DNA extraction Protocol | Recommended for extracting good quality, high molecular weight (HMW) DNA (> 80 kb). Substituting other extraction kits may adversely affect system performance and are not supported.
Eppendorf | Centrifuge 5424, with aerosol-tight Rotor FA-45-24-11** | 022620401 | RT | DNA extraction Protocol | Need a microcentrifuge that reaches 15,000 X g to pellet cells. Other models ok to use here.
Thermo Fisher Scientific | Gibco PBS, pH 7.4 ** | 10010-023 | RT | DNA extraction Protocol | Recommended for washing cells. Other PBS is ok here.
| DynaMag™-2 Magnet Rack* | 12321D | RT | DNA extraction Protocol | Required for washing steps in extraction protocol. If set up for WES libraries, material is also required.
^ Note that other items are required for Training, and Whole Genome or Whole Exome Sequencing after training. Refer to the Chromium User Guide for complete workflow requirements.
* Items in bold red have been validated by 10X Genomics and are required for Chromium workflow training and system operation. No substitutions are allowed.
** Listed materials are compatible with workflow. Substituting materials may adversely affect system performance and are not supported.
*** Example materials listed. Substituions may be used.

## Automation
Source | Description | Part # | Storage Conditions | Step(s) in protocol | Why needed
10x | 10x Custom Chip Holder Fixture | Custom | RT | GEM Generation and Barcoding | Supports Chip Holders during chip loading and unloading with unique positioning for left and right chips
| 10x Custom Chip Holder Left | Custom | RT | GEM Generation and Barcoding | Custom Chip Holder with additional unique hole Left of the back label, for use with the Chip Holder Fixture
| 10x Custom Chip Holder Right | Custom | RT | GEM Generation and Barcoding | Custom Chip Holder with additional unique hole Right of the back label, for use with the Chip Holder Fixture
| 10x Custom Plate Holder (2) | Custom | RT | Throughout Protocol | Securely hold plates during processing (critical during foil piercing) and maintains good thermal contact with the deck
Agilent | 250uL Filtered, 96 Tips/Rack, Sterile | 19477-022 | RT | Throughout Protocol | Automation tips compatible with emulsion and preserving DNA length
| PCR Plate Insert | G5498B#013 | RT | Throughout Protocol | Efficient labware for commonly-used reagents
| Deepwell Plate Insert | G5498B#012 | RT | Throughout Protocol | Efficient labware for commonly-used reagents
| Seahorse single cavity polypropylene reservoir | 201254-100 | RT | Throughout Protocol | Efficient labware for commonly-used reagents
| Seahorse 12 column polypropylene reservoir | 201256-100 | RT | Throughout Protocol | Efficient labware for commonly-used reagents
| Seahorse 24 column polypropylene reservoir | 201296-100 | RT | Bead-based Cleanups | Efficient labware for commonly-used reagents
Alpaqua | Magnum FLX Enhanced Universal Magnet Plate* | A000400 | RT | All bead-based pull-downs | Quickly pellets beads at bottom of tube, enabling efficient low-volume elution and enhanced resuspension
Thermo Fisher Scientific | 96 DeepWell PP Plate Natural RNAse/DNAse-Free, Sterile, 1 mL | 260251 | RT | Make Plate creation | Accepts higher-volume, singly-column input for distribution across multiple columns in normal 96-well plate
| Quant-iT PicoGreen dsDNA Assay Kit* | P7589 | Various | Quantitation steps | Accurate and scalable quantitation method
Sigma | Sodium hydroxide solution | S2770-100ML | RT | Reagent Setup | 1.0 N, BioReagent, suitable for cell culture
VWR | HSW Norm-Ject Sterile Luer-Lock Syringes, Air-Tite* | 80076-426 | RT | Reagent Setup | Needed for preparing Denaturing Reagent
| Syringe Filters with Polypropylene Housing* | 28145-483 | RT | Reagent Setup | Needed for preparing Denaturing Reagent
* Items in bold red have been validated by 10X Genomics and are required for Chromium workflow training and system operation. No substitutions are allowed.

## WGS additional
Source | Description | Part # | Storage Conditions | Step(s) in protocol | Why needed
Bio-Rad | Hard-Shell® High Profile 96-Well Semi-Skirted PCR Plates *** | HSS-9601 | RT | Library Quantitation | Used for serial dilutions for qPCR - any plate will do.
| Hard-Shell Low-Profile Thin-Wall 96-Well Skirted PCR Plates*** | HSP-9665 | RT | Library Quantitation | Required for qPCR.
| CFX Connect real-time PCR Detection System*** |  | RT | Library Quantitation | Any Real-Time Detection system will do.
KAPA | Illumina Library Quantification Kit* | KK4824 | -20°C | Preparing Master Mix | Required for library quant steps
Eppendorf | Centrifuge 5430, with Rotor FA-45-30-11 incl. rotor lid, keypad, 120 V/50 – 60 Hz** | 022620509 | RT | Plating the Master Mix and Samples | Need a centrifuge for spinning plates to 130 g. Other models available.
* Items in bold red have been validated by 10X Genomics and are required for Chromium workflow training and system operation. No substitutions are allowed.
** Listed materials are compatible with workflow. Substituting materials may adversely affect system performance and are not supported.
*** Example materials listed. Substituions may be used.

## WES-Agilent Only
Source | Description | Part # | Storage Conditions | Step(s) in protocol | Why needed
Agilent | SureSelectXT Human All Exon V6 |  | -80°C | Preparing Capture Baits | Capture baits. Required
| SureSelectXT Reagent Kit, HSQ | G9611B | RT/-20°C | Preparing Hybridization Buffer, Preparing the Blocking Mix, Preparing Reagent Loading and Plating | Capture additional reagents. Required
Eppendorf | Vacufuge® Plus** | 022820001 | RT | Library Preparation | Needed to dry down sample
Torrey Pines Scientific | Block, 30 x 1.5 ml (optional) | 620-5013 | RT | Pre-warming SureSelect Wash 2 Buffer | Can be used for heating small volumes of wash buffer 2
| Block, 12 x 15 ml (optional) | 620-5027 | RT | Pre-warming SureSelect Wash 2 Buffer | Can be used for large volume heating of wash buffer 2
| Echotherm™ IC20 Digital, Electronic Chilling/Heating Dry Bath (optional) | IC20 | RT | Pre-warming SureSelect Wash 2 Buffer | Needed for use with Torrey Pines Heat blocks
VWR | Conical Bottom Centrifuge Tube with Flat Cap, 50mL*** | 89039-666 | RT | Exome Target Enrichment | Can be used for heating large volume wash buffer 2
GAST | DOA Diaphragm Pump** | DOA-P704-AA | RT | Exome Target Enrichment | Required for Vacufuge
IDT | Universal Blocking Oligo - TS-p5 | 1016184 | lyophilized RT, resuspended 4°C | Exome Target Enrichment | Required blocking oligo for 10X library
| Universal Blocking Oligo - TS-p7(8nt) | 1016188 | lyophilized RT, resuspended 4°C | Exome Target Enrichment | Required blocking oligo for 10X library
Thermo Fisher Scientific | DynaBeads® MyOne™ Streptavidin T1* | 65601 | 4°C | Preparing Streptavidin Magnetic Beads , Binding the Capture Library to Streptavidin Magnetic Beads | Required for capturing molecules bound to baits.
| MicroAmp® Optical Film Compression Pad* | 4312639 | RT | Preparing Reagent Loading and Plating | Required to minimize evaporation during overnight hyb
| DynaMag™-2 Magnet Rack* | 12321D | RT |  | Required for washing capture beads
Bio-Rad | Tube Frame | 1849001 | RT | Exome Target Enrichment | Required for preventing crushing of tubes. This part is initially supplied with the C1000.
| Microseal 'B' Adhesive Seals** | MSB-1001 | RT | Exome Target Enrichment | Required to minimize evaportaiton during overnight hyb. Other sticky seals may lead to sample loss and reduced performance.
Covaris | microTUBE AFA Fiber Screw-Cap 6 x 16 mm** | 520096 | RT | Shearing | Required for shearing with Covaris M220
| M220 Focused-Ultrasonicator™** | 500295 | RT | Shearing | Required for shearing DNA. Only Covaris sonicators recommended.
| M220 Holder microTUBE** | 500301 | RT | Shearing | Required for shearing with microTUBE
* Items in bold red have been validated by 10X Genomics and are required for Chromium workflow training and system operation. No substitutions are allowed.
** Listed materials are compatible with workflow. Substituting materials may adversely affect system performance and are not supported.
*** Example materials listed. Substituions may be used.